## Audit QA Results
# | Audit Question | Sub-queries Used | Answer | Reference Files Used | Chunks Used
1
2 | What criteria are used to classify a product as a medical device under applicable regulatory jurisdictions? | regulatory criteria for medical device classification jurisdiction-specific guidelines for medical device categorization rules and standards for classifying products as medical devices | The criteria used to classify a product as a medical device under applicable regulatory jurisdictions for the Halcyon 4.0 (HAL 4.0) radiotherapy system are detailed as follows: 1. **United States**: In the United States, HAL 4.0 is classified as a Class II medical device under 21 CFR §892.5050, which corresponds to the product code IYE - Accelerator System, Linear [Regulatory / Regulatory Plan / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device]. 2. **European Union**: HAL 4.0 is classified as Class IIb under the EU Medical Device Regulation 2017/745 (MDR), specifically applying Annex VIII Classification Criteria, Rule 9 [Regulatory / Regulatory Plan / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device]. 3. **Canada**: In Canada, HAL 4.0 is classified as a Class III medical device according to the Medical Device Regulations (SOR/98-282), Schedule 1, Rule 8 [Regulatory / Regulatory Plan / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device]. 4. **Japan**: HAL 4.0 is classified as a Class III device under the Pharmaceuticals and Medical Devices Act, Article 2, Paragraph 5 [Regulatory / Regulatory Plan / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device]. 5. **China**: The device falls under Class III according to the National Medical Products Administration (NMPA) medical device classification category, specifically identified under 05-01-01 for medical electron accelerators [Regulatory / Regulatory Plan / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device]. The classification criteria involve assessing the device's risk, use, and regulatory requirements unique to each region. Each region has its framework for classifying medical devices based on their intended use and risk assessment. Documents cited: - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device | Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page 5, Section: 4. Type of Device | 200
3 | Are all devices, including accessories, software, variants, and families, appropriately identified and included? | device accessory and component identification list software version tracking and configuration documentation list of device variants and product family classifications | The following documents provide relevant evidence regarding whether all devices, including accessories, software, variants, and families, are appropriately identified and included for the Halcyon 4.0 radiotherapy system: 1. **Accessories and Components**: - The document titled "HAL 4.0 Accessory and Component List" clearly lays out the accessories and components included with the HAL 4.0 device. It lists specific components necessary for system functionality, such as lasers, transformers, main disconnect panels, and relay junction boxes [Design Output Specification / Design Descriptions / D46704 HAL 4.0 Accessory and Component List Rev 1.0.docx.pdf, page 4-6]. Additionally, it mentions no clinical-use accessories are required, which provides a direct statement about the necessity or optional nature of additional accessories [Design Output Specification / Design Descriptions / D46704 HAL 4.0 Accessory and Component List Rev 1.0.docx.pdf, page 3]. 2. **Device Models and Families**: - The document describes the HAL 4.0 platform as supporting two product models: Halcyon and Ethos, which are designated as medical linear accelerators with certain capabilities. Ethos includes additional capabilities for adaptive treatment solutions [Regulatory / Test Reports - Regulatory Compliance / 102341812MIN-003a - HAL4.0 - IEC 60601-1-3.pdf, page 8-9]. This differentiation ensures that different models under the HAL platform are identified and included. 3. **Software and Configuration Items**: - Software configuration is detailed with specific identifiers for software versions used on the HAL 4.0 system. Unique identifiers ensure traceability and management of configuration items, signifying the rigorous control over software elements integrated into the system [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, page 286-287]. Moreover, a specific list of software tools and system configurations utilized in tests is documented, supporting comprehensive coverage of software components [Design Validation / Test Environment / D77635 Halcyon SN1778 TSCL Rev 14.pdf, page 9-10]. The available excerpts establish that the Halcyon 4.0 radiotherapy system has its accessories, components, models, and software variants appropriately identified and documented as part of the device's evidence of compliance. However, there is no explicit evidence that lays out how every potential device element, especially rare or optional components, is captured or traced in broader implementation contexts beyond what is described. Documents cited: - Design Output Specification / Design Descriptions / D46704 HAL 4.0 Accessory and Component List Rev 1.0.docx.pdf, page 3-6 - Regulatory / Test Reports - Regulatory Compliance / 102341812MIN-003a - HAL4.0 - IEC 60601-1-3.pdf, page 8-9 - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, page 286-287 - Design Validation / Test Environment / D77635 Halcyon SN1778 TSCL Rev 14.pdf, page 9-10 | Design Output Specification / Design Descriptions / D46704 HAL 4.0 Accessory and Component List Rev 1.0.docx.pdf, page 3-6 Regulatory / Test Reports - Regulatory Compliance / 102341812MIN-003a - HAL4.0 - IEC 60601-1-3.pdf, page 8-9 Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, page 286-287 Design Validation / Test Environment / D77635 Halcyon SN1778 TSCL Rev 14.pdf, page 9-10 | 200
4 | How are changes such as maintenance releases, configuration changes, or software updates evaluated to determine if design controls apply? | evaluation process for maintenance release impact on design controls assessment of configuration changes for design control requirements software update evaluation in relation to design controls | The evaluation of changes such as maintenance releases, configuration changes, or software updates to determine if design controls apply is evidenced by several documents related to HAL 4.0. Here's a summary of the provided excerpts: 1. **Maintenance Release Documents**: - The HAL 4.0 MR1 Maintenance Release Document specifies that no hardware changes are involved, and the release scope is limited to software changes only, which have defined regulatory and project management scopes without introducing new functionality or impacting existing device interfaces [Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, page 10]. - For both MR1 and MR2 maintenance releases, there is an explicit statement that these changes do not introduce new risks or hazards, and the risk management reports do not require updates [Design Changes / Maintenance Releases / D95424 HAL 4.0 MR2 (HAL Evolve) Maintenance Release Document Rev 04.pdf, pages 25-26; Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, pages 27-28]. 2. **Project Management and Design Controls**: - In the documents, it is noted that no separate Design Output Review will be conducted for certain projects. Instead, an approval of the final maintenance release document confirms the closure of design output activities [Design Changes / Maintenance Releases / D95424 HAL 4.0 MR2 (HAL Evolve) Maintenance Release Document Rev 04.pdf, page 9-10; Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, page 12]. 3. **Risk Management and Configuration Management**: - Necessary design changes are to be handled via change request, and a risk assessment is to be performed for any update necessitating design changes. This includes documenting any test activities and assessing the scope of updates [Regulatory / Test Reports - Regulatory Compliance / D78333 Ethos 4.0 Description of interoperability Study Rev01.docx.pdf, page 21]. - Configuration management for software is managed per specified procedures, ensuring that changes are controlled and follow established processes, such as impact analysis and verification & validation tasks [Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, pages 27-28]. 4. **Change Request Process**: - Multiple change request documents outline the criteria necessitating a change request. Major changes in functionality, design, or compliance thresholds trigger the mandatory creation of a change request to assess the impact and determine the necessary approval process [Design Changes / Change Requests / D52222 HAL 4.0 Change Request 2 rev 01.pdf, page 1; Design Changes / Change Requests / D64997 HAL4.0 Change Request_Kemnath takes place of Crawley to build HAL4.0 Rev 01.pdf, page 1]. Overall, these documents demonstrate a structured approach to evaluate changes through defined project management processes, risk assessment, configuration management, and change request protocols to determine when design controls are applicable for HAL 4.0 system changes. **Documents Cited:** - Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, page 10 - Design Changes / Maintenance Releases / D95424 HAL 4.0 MR2 (HAL Evolve) Maintenance Release Document Rev 04.pdf, pages 25-26 - Design Changes / Maintenance Releases / D95424 HAL 4.0 MR2 (HAL Evolve) Maintenance Release Document Rev 04.pdf, page 9-10 - Regulatory / Test Reports - Regulatory Compliance / D78333 Ethos 4.0 Description of interoperability Study Rev01.docx.pdf, page 21 - Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, pages 27-28 - Design Changes / Change Requests / D52222 HAL 4.0 Change Request 2 rev 01.pdf, page 1 - Design Changes / Change Requests / D64997 HAL4.0 Change Request_Kemnath takes place of Crawley to build HAL4.0 Rev 01.pdf, page 1 |  | 200
5 | Are legacy devices assessed for compliance with current design control requirements? | legacy device compliance assessment report for current design controls evaluation of design control updates for legacy devices legacy device redesign and compliance strategy documentation | The retrieved document excerpts do not provide direct information or evidence regarding the assessment of legacy devices for compliance with current design control requirements specific to the Halcyon 4.0 (HAL 4.0) radiotherapy system. The excerpts focus mainly on risk control measures, validation strategies, and testing outcomes related to the device but do not specifically address the compliance assessment of legacy devices under current design control requirements. Documents cited: - None (no relevant evidence found) | None (no relevant evidence found) | 200
6 | How does the organization ensure global regulatory requirements are considered in determining design control applicability? | process for integrating global regulatory requirements into design controls criteria for determining design control applicability based on international standards documentation of compliance with global regulatory requirements during the design phase | The excerpts provide several insights into how the organization ensures global regulatory requirements are addressed in determining design control applicability for the Halcyon 4.0 radiotherapy system. Key documents and procedures are utilized to comply with global regulatory requirements: 1. **Varian's Design and Development Procedure (DND01):** - The organization employs a process based on the Rational Unified Process, detailed in document DND01 Design and Development, which outlines project phases, documentation requirements, and management oversight needed to comply with global regulatory requirements, specifically focusing on US FDA 21 CFR 820 and ISO 13485 [Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page: 23]. 2. **Risk Management:** - The RSK01 Medical Device Risk Management procedure supplements the DND01 and aligns with EN ISO 14971:2019, addressing risk management implementation across all phases of development [Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page: 23]. - The risk control measures are implemented and verified to reduce or maintain acceptable risk levels, utilizing international standards as referenced in the Risk Management Plan [Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf, Page: 9]. 3. **Regulatory Plan (D54978):** - It is the responsibility of the project team to ensure all regulatory requirements, including national language requirements, are included within the design and development plans according to current Good Manufacturing Practices (cGMPs) and national legislation, with Section 8 specifically containing country-specific regulatory standards [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 9]. 4. **Applicable Standards Documentation:** - The document D45070 Hal 4.0 Applicable Product Standards evaluates the applicability of various product standards necessary for market access globally, specifying project-specific requirements in the Regulatory Plan and Design and Development Plan [Regulatory / Applicable Standards / D45070 Hal 4 Applicable Product Standards Rev 2.docx.pdf, Page: 3-4]. 5. **Quality Systems:** - Compliance with EN ISO 13485 ensures the quality management process is aligned with international standards, helping meet regulatory criteria such as those in the MDR GSPR, Canadian MDR, and the Australian Essential Principles [Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf, Page: 44-45]. These documents and processes demonstrate that Varian's approach to design control applicability considers and integrates global regulatory requirements into their product development and management systems, as evidenced in these explicit references to standards and processes. Documents cited: - Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page: 23 - Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf, Page: 9 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 9 - Regulatory / Applicable Standards / D45070 Hal 4 Applicable Product Standards Rev 2.docx.pdf, Page: 3-4 - Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf, Page: 44-45 | Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page: 23 Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf, Page: 9 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 9 Regulatory / Applicable Standards / D45070 Hal 4 Applicable Product Standards Rev 2.docx.pdf, Page: 3-4 Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf, Page: 44-45 | 200
7 | Does the organization have documented design and development procedures compliant with applicable regulations? | design and development procedures documentation regulatory compliance for design and development evidence of adherence to design and development regulations | The organization has documented design and development procedures that show compliance with applicable regulations for the HAL 4.0 radiotherapy system. The procedures are clearly outlined in various documents which detail the standards and processes followed during the design and development phases: 1. **Design and Development Procedure (DND01):** This is a key document that defines the phases of a project, including documentation requirements for each phase. It's designed to comply with global regulatory requirements, specifically US FDA 21 CFR 820 and ISO 13485 [Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page 23]. 2. **Design and Development Plan (DND04T01):** This plan references and integrates the DND01 procedure and includes detailed planning for software development, risk management, and configuration management. It is part of the deliverables of the Inception and Elaboration Phases [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 226-227]. 3. **Software Development Planning:** The manufacturer has established procedures to cover various aspects such as software safety classifications, traceability between system requirements, risk control measures, and the coordination of software and system development [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 223-224]. 4. **Configuration Management Procedures (DND19):** These procedures link responsibilities to tasks and include activities and tasks for configuration management, ensuring the DND04T01 Design and Development Plan documents the controlled items [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 239-240]. 5. **Regulatory Plans and Standards Compliance:** Various documents, such as the Regulatory Plan and the Applicable Product Standards Record, demonstrate the product's compliance with required safety standards, including IEC 60601-1 and EN ISO 13485, ensuring that the development process aligns with established international regulations [Regulatory / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page 9; Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, Page 4]. These documents indicate that the organization's procedures align with regulatory standards, suggesting a structured and compliant approach to the design and development of the HAL 4.0 radiotherapy system. Documents cited: - Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page 23 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 226-227 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 223-224 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 239-240 - Regulatory / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page 9 - Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, Page 4 | Regulatory / D70036 HAL 4.0 Software Description Report China Submission Rev 03 .docx.pdf, Page 23 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 226-227 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 223-224 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 239-240 Regulatory / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page 9 Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, Page 4 | 200
8 | Are these procedures consistently applied to all identified devices? | device-specific procedure application records consistency validation reports for device procedures procedure adherence monitoring across all device categories | The retrieved excerpts provide detailed information on the procedures followed by the Halcyon 4.0 radiotherapy system, but they do not explicitly address whether these procedures are consistently applied across all identified devices. The excerpts describe various procedural aspects related to risk management, design, verification, and validation for the Halcyon 4.0, demonstrating compliance to IEC and ISO standards and internal procedures. However, there is no direct information confirming that these procedures are consistently applied to all devices other than Halcyon 4.0. The design and development procedures (DND01, DND04T01, etc.) and risk management procedures (RSK0101) cited throughout the sources are primarily associated with Halcyon 4.0. Furthermore, the device records and test reports specifically mention Halcyon 4.0, indicating that the information pertains explicitly to this device (e.g., [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf], [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf]). Thus, while the documents contain detailed procedures and evidence of their application to Halcyon 4.0, there is no evidence from the retrieved excerpts indicating these procedures are consistently applied to other identified devices. Documents cited: - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf | Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf | 200
9 | How does the organization ensure design controls are applied across the lifecycle for new, modified, and legacy devices? | design control procedures for new device development modification approval process for device changes legacy device design control documentation and compliance | The provided document excerpts offer detailed insights into the organization's approach to ensuring design controls are applied across the lifecycle for new, modified, and legacy devices. Here is an overview of how these controls are managed for the Halcyon 4.0: 1. **Design and Development Procedures**: The organization uses the DND01 Design and Development Procedure to manage design controls, as referenced in multiple documents [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 222-223; D101277 106485676MIN-001, Page: 290-291]. This procedure includes the software development lifecycle, configuration management, and tools utilized in developing the medical device software [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001, Page: 240-242]. 2. **Risk Management**: Risk management is a consistent part of the lifecycle management. The RSK01 Medical Device Risk Management procedure is used in conjunction with the DND01 Design and Development Procedure to evaluate risks during all phases of development, as well as the management of residual risks [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 292-293]. 3. **Traceability and Verification**: The design and development process includes traceability from system requirements to risk control measures, ensuring that all changes are documented and verified. This is supported by maintaining design and development plans and trace matrices [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 282-283; D68599 Hal 4.0 Design Verification Trace Matrix]. 4. **Configuration Management**: Software configuration management is documented in the DND04T01 Design and Development Plan, ensuring all configuration items are properly identified and controlled throughout the lifecycle [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 287-288]. 5. **Change Management**: Changes to the software and configuration items are managed through a formal change request process, documented in DND04F01 Change Request and related procedural documents to ensure all changes are verified and approved [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 293-294]. 6. **Verification and Validation**: The development process includes rigorous verification and validation activities to ensure all design outputs meet established requirements and any failures are effectively addressed [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001, Page: 226-227; 104993953MIN-001, Page: 278-279]. These processes delineate how the organization systematically applies design control procedures to maintain compliance and manage risks across the lifecycle for new, modified, and legacy devices such as the Halcyon 4.0. Documents cited: - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf - Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf | Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf | 200
10 | Are roles and responsibilities for design activities clearly defined? | design roles and responsibilities matrix design team roles and accountability chart responsibility assignment for design activities | The roles and responsibilities for design activities for the Halcyon 4.0 radiotherapy system appear to be clearly defined across several documents: 1. **User Assistance Plan**: The "RACI for Design & Development" outlines key responsibilities and roles in developing the Instructions for Use, involving subject-matter experts in content development and formal review processes [Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 7]. 2. **Marketing Claims Library**: Responsibilities for design activities related to marketing claims are specified, including roles such as Product Marketing Manager, Product Manager, SW QA Engineer, and PRC Regulatory, each assigned specific tasks [Design Input Specification / Product Requirement Documents / D53708 Marketing Claims Library HAL 4.0 Rev 4.0.docx.pdf, Pages: 4-5]. 3. **Maintenance Release Document**: This document specifies roles in design implementation, project management, and product management for the HAL 4.0 maintenance release, outlining the involvement of various team members [Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, Pages: 7-8]. 4. **Risk Management Report**: Assigns specific responsibilities and authorities to roles such as Risk Manager, Quality Assurance, Regulatory Affairs, and various engineering disciplines, highlighting a structured approach to risk management activities [Risk Management / Risk Management Reports / D56979 HAL 4.0 Risk Management Report rev07.pdf, Page: 5]. 5. **Design and Development Plan**: Details the team structure and roles, including responsibilities for project management, systems engineering, and product management, supporting the overall design and development process [01. Design and Development Planning / 01. System / D23997 Halcyon v2 Design and Development Plan Rev 01.pdf, Pages: 9, 37-39]. 6. **Traceability Plan**: Assigns tasks related to traceability evaluation and identifies responsibilities for regulatory and quality assurance aspects [Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf, Pages: 5-6]. These documents collectively outline a comprehensive assignment of responsibilities for design-related activities, ensuring clear roles are defined within the development lifecycle of the Halcyon 4.0 system. Documents cited: - Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 7 - Design Input Specification / Product Requirement Documents / D53708 Marketing Claims Library HAL 4.0 Rev 4.0.docx.pdf, Pages: 4-5 - Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, Pages: 7-8 - Risk Management / Risk Management Reports / D56979 HAL 4.0 Risk Management Report rev07.pdf, Page: 5 - 01. Design and Development Planning / 01. System / D23997 Halcyon v2 Design and Development Plan Rev 01.pdf, Pages: 9, 37-39 - Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf, Pages: 5-6 | Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 7 Design Input Specification / Product Requirement Documents / D53708 Marketing Claims Library HAL 4.0 Rev 4.0.docx.pdf, Pages: 4-5 Design Changes / Maintenance Releases / D69533 HAL 4.0 MR1 Maintenance Release Document Rev09.docx.pdf, Pages: 7-8 Risk Management / Risk Management Reports / D56979 HAL 4.0 Risk Management Report rev07.pdf, Page: 5 01. Design and Development Planning / 01. System / D23997 Halcyon v2 Design and Development Plan Rev 01.pdf, Pages: 9, 37-39 Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf, Pages: 5-6 | 200
11 | How is risk integrated into design and development activities? | risk analysis and management process during device design design controls and risk mitigation strategies in development documentation of risk assessment and integration in design phases | The integration of risk into design and development activities for the HAL 4.0 device follows a structured risk management process conforming to ISO 14971, a standard for risk management of medical devices. This process is integrated at various stages of the device lifecycle and documented in multiple reports and plans. Here are specific ways risk is integrated into the design and development activities: 1. **Risk Management System and Iterative Process**: HAL 4.0 follows a continuous iterative risk management process throughout the device lifecycle. This involves establishing, implementing, documenting, and maintaining a risk management system. Risk management activities include the creation of risk management plans for each device, identification and analysis of hazards, and re-evaluation of risks when software requirements are established [Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf ¦ Page: 3-4], [Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf ¦ Page: 4]. 2. **Risk Management Plan and Activities**: A specific risk management plan (RMP) for HAL 4.0 is in place, detailing the activities and processes to manage risks. This includes conducting top-down risk hazard analysis, bottom-up risk failure mode and effects analysis (FMEA), and regular reviews by competent personnel [Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 6], [Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 7]. 3. **Verification and Validation**: Risk control measures identified and implemented during the design and development are verified and validated. This is documented in the Risk Management Report and other verification and validation reports, which confirm the implementation and effectiveness of risk control measures [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 292-293]. 4. **Post-production Monitoring**: Production and post-production information are continuously monitored to identify new risks or reassess known risks. This ongoing process informs updates to risk management activities based on new information from production changes, defect identifications, and external reports [Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 9-10]. 5. **Integration into Design Control**: Risk control measures are integrated into system and software requirements through a design and development procedure. This ensures a systematic connection between risk controls and software functionality, complying with ISO 14971 [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf ¦ Page: 242-243]. Overall, the HAL 4.0's approach to integrating risk into design and development is comprehensive, involving systematic risk assessment processes, continuous monitoring, and verification of implemented controls across all device lifecycle phases. Documents cited: - Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf ¦ Page: 3-4 - Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf ¦ Page: 4 - Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 6 - Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 7 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 292-293 - Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 9-10 | Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf ¦ Page: 3-4 Regulatory / International Submission Dossier / D70060 HAL 4.0 GSPR Checklist Rev2.docx.pdf ¦ Page: 4 Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 6 Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 7 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 292-293 Risk Management / Risk Management Plans / D59555 HAL 4.0 Risk Management Plan rev03.pdf ¦ Page: 9-10 | 200
12 | Is technical documentation established and maintained for each device? | technical documentation creation and maintenance procedures device-specific technical file documentation recordkeeping for technical documentation updates | The available excerpts confirm that technical documentation is established and maintained for the Halcyon 4.0 device. 1. **System Requirements Specification (SyRS):** This document includes installation and maintenance considerations specific to the device, indicating that technical requirements are clearly documented [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 248-249]. 2. **Design History File:** A comprehensive Design History File Index is maintained, listing all required project documents, their titles, identification numbers, revisions, and storage locations. This documentation is referenced within the document control system [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 13-14]. 3. **User Needs and Workflows:** Documentation addresses user needs with clear and accurate customer-facing instructions on performing tasks, including maintenance, suggesting that operational documentation for users is in place [Design Input Specification / Product Requirement Documents / D57043 HAL 4.0 User Needs and Detailed Workflows Rev 2.0.docx.pdf ¦ Page: 67-68]. 4. **Service Documentation:** Technical documentation provides written procedures for maintenance and servicing tasks, indicating that service instructions are established and maintained for the Halcyon 4.0 [Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf ¦ Page: 68]. 5. **Compliance with Standards:** The technical documentation is set up to meet regulatory and compliance standards, including maintaining records for verification and risk management, as presented in the compliance reports [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 249-250]. 6. **Service User Needs:** Documentation supporting service activities includes detailed workflows and instructions for the Customer Support Representative, ensuring that procedures are documented for various scenarios encountered during maintenance [Design Input Specification / Product Requirement Documents / D62523 HAL 4.0 Service User Needs and Detailed Workflows Rev 3.0.docx.pdf ¦ Page: 6]. 7. **Regulatory Documentation and Strategy:** The Halcyon 2.0 Regulatory Plan outlines documentation requirements for various markets, which likely extend to Halcyon 4.0, ensuring global compliance [01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf ¦ Page: 61-62]. These documents confirm that technical documentation is not only established but also continuously updated and managed to support the device's requirements across installations, maintenance, user needs, and regulatory compliance. Documents cited: - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 248-249 - Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 13-14 - Design Input Specification / Product Requirement Documents / D57043 HAL 4.0 User Needs and Detailed Workflows Rev 2.0.docx.pdf ¦ Page: 67-68 - Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf ¦ Page: 68 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 249-250 - Design Input Specification / Product Requirement Documents / D62523 HAL 4.0 Service User Needs and Detailed Workflows Rev 3.0.docx.pdf ¦ Page: 6 - 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf ¦ Page: 61-62 | Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 248-249 Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 13-14 Design Input Specification / Product Requirement Documents / D57043 HAL 4.0 User Needs and Detailed Workflows Rev 2.0.docx.pdf ¦ Page: 67-68 Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf ¦ Page: 68 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ Page: 249-250 Design Input Specification / Product Requirement Documents / D62523 HAL 4.0 Service User Needs and Detailed Workflows Rev 3.0.docx.pdf ¦ Page: 6 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf ¦ Page: 61-62 | 200
13 | Does the documentation meet regulatory requirements for applicable markets? | regulatory approval documents and certifications for primary markets compliance checklist and market-specific requirements labeling, instructions, and regulatory documentation for distribution | The documentation partially meets regulatory requirements for applicable markets. The documents retrieved provide some evidence of compliance with standards and documentation requirements for specific markets, but comprehensive evidence of compliance across all applicable standards and markets is incomplete based on the retrieved excerpts. 1. **EU Requirements:** - Conformity with EU MDR is demonstrated through compliance with EN standards, such as IEC 62304 for medical device software and IEC 62366-1 for usability engineering [Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, page 7]. - The documentation for the EU includes requirements related to the CE Mark and International Submission Dossier (ISD), indicating that certain documents must be submitted electronically and comply with evolving EU regulatory requirements [01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, page 50]. 2. **Brazil Requirements:** - Brazil requires documentation of quality certificates, software versions, safety and efficacy data, and test reports conforming to IEC standards [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 64-65]. Anvisa certification of conformity is required if an IEC standard is translated for Brazilian use [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 66-67]. - Key documents include the INMETRO certification and Good Manufacturing Practice (GMP) certificates approved by ANVISA [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, page 64-65]. 3. **Canada Requirements:** - For Health Canada, the document outlines the need for a medical device license application and compliance with ISO 13485 [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, page 45]. - Additional requirements include documentation for device descriptions, safety and effectiveness information, and complaint and recall handling [01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, pages 54-55]. 4. **Korea, China, and Japan Requirements:** - Documentation indicates compliance with country-specific standards, such as IEC 60601 standards for Korea and product labeling requirements for China [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, page 69]. - The Shonin process in Japan requires a range of documents, including sales data, schematics, and Japanese labeling [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 48-49]. The documentation appears to address various regulatory strategies. However, there is a lack of specific certification outcomes, comprehensive test reports, or declarations of conformity provided in the excerpts for all applicable standards and markets, such as the requirements specific to the US FDA or the broader set of countries within the Medical Device Single Audit Program (MDSAP). Documents cited: - 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, pages 50, 54-55 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 45, 64-67, 69 - Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, page 7 | 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, pages 50, 54-55 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 45, 64-67, 69 Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, page 7 | 200
14 | Is there a defined structure for technical documentation, such as DHF, STED, or Technical File? | design history file (DHF) documentation structure summary technical documentation (STED) format and organization technical file structure and required contents | The excerpts provided offer evidence that there is a defined structure for technical documentation regarding the Design History File (DHF) for the Halcyon 4.0 radiotherapy system. Key points from the documentation include: 1. **Design History File Index (DHFI):** The Design History File Index for Halcyon 4.0 explicitly lists all required project documentation for the released version of the device, including titles, identification numbers, revision numbers, and storage locations (either physical or electronic) [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14, Section: 1. Introduction > 1.1 Scope]. This indicates a structured approach to document organization and retrieval. 2. **Documentation Structure:** The DHFI includes various sections that outline different aspects of the design and development process, such as Design and Development Planning, Design Input Specification, Design Output Specification, Design Reviews, Design Verification, Design Validation, and Design Transfer, among others [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Pages 13-18, Section: 2. Design History File Index]. This compartmentalization suggests a methodical framework for managing technical documents. 3. **Verification and Validation Records:** The DHF repository maintains records of testing, including verification and validation test records, which are part of the electronic DHF repository. This is reinforced by the retrieval and documentation of verification results, as evidenced by multiple excerpts referencing the DHF storage of such information [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages 309-310, Section: 4 Supply voltage: (Continued)]. 4. **Regulatory Documentation:** Additional documents cited include the Risk Management Report and Validation Summary Reports, which are integral parts of the regulatory framework and provide detailed insights into the risk assessment and testing processes associated with the device design [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page 265-266, Section: 4 Supply voltage: (Continued)]. In conclusion, the documents retrieved clearly outline a defined structure for technical documentation pertaining to Halcyon 4.0, incorporating various plans, reports, and procedural instructions encompassed within the DHF, which assures compliance and traceability in the device development and regulatory process. Documents cited: - Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14, Section: 1. Introduction > 1.1 Scope - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages 309-310, Section: 4 Supply voltage: (Continued) | Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14, Section: 1. Introduction > 1.1 Scope Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages 309-310, Section: 4 Supply voltage: (Continued) | 200
15 | How is completeness and consistency of technical documentation ensured? | process for ensuring completeness in technical documentation methods to verify consistency in technical documentation quality control procedures for reviewing technical documents | The completeness and consistency of technical documentation for the Halcyon 4.0 radiotherapy system are ensured through a variety of processes and reviews outlined in several documents: 1. **Design Output Review (DOR):** The DND07F06 Design Output Review (DOR) Rev 01 ensures completeness of all documentation, activities, and tasks, and the Halcyon 4.0 Design Output Review Rev 01 further assures that these elements are complete [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page: 281-282]. 2. **Quality Control and Reviews:** Internal and formal reviews are conducted to assure the accuracy, completeness, and quality of user assistance content, as described in several documents. Peer and editorial reviews are undertaken to ensure style, clarity, consistency, and other quality aspects [Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 5; Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf, Page: 3-4; Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page: 2-4]. 3. **Validation Processes:** Validation completion criteria include ensuring that all activities defined in the Validation Plan have been performed and documented in the approved Validation Summary Report. Acceptance criteria require reviewing and prioritizing open defects and ensuring no safety issues exist, quality targets are met, and essential documents are approved [Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf, Page: 32; Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf, Page: 16-17]. 4. **Verification Summary Reports:** These reports document procedures, environments used, and ensure that residual anomalies are evaluated to avoid unacceptable risks [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Page: 279-280]. 5. **Packaging Validation:** Packaging documentation is verified to ensure it maintains integrity under specified conditions, such as transport and storage [Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf, Page: 71; Design Validation / Test Reports / D65245 HAL 4.0 Packaging Validation Protocol and Report Rev 04.docx.pdf, Page: 3]. 6. **Documentation Traceability:** The technical documentation is supposed to be free from contradictions within itself, validated for understandability, and arranged in a verifiable manner as part of the documentation review process [Regulatory / Test Reports - Regulatory Compliance / D75417 IEC 25051-GB T25000 Compliance for HAL 4.0 Rev 01.pdf, Page: 17-18; Page: 25]. These processes and policies form the basis for ensuring that the Halcyon 4.0's technical documentation is complete and consistent across its lifecycle. Documents cited: - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf - Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf - Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf - Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf - Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf - Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf - Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf - Design Validation / Test Reports / D65245 HAL 4.0 Packaging Validation Protocol and Report Rev 04.docx.pdf - Regulatory / Test Reports - Regulatory Compliance / D75417 IEC 25051-GB T25000 Compliance for HAL 4.0 Rev 01.pdf | Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf Design Validation / Test Reports / D69032 HAL 4.0 Validation Test Log Rev 01.pdf Design Validation / Test Reports / D65245 HAL 4.0 Packaging Validation Protocol and Report Rev 04.docx.pdf Regulatory / Test Reports - Regulatory Compliance / D75417 IEC 25051-GB T25000 Compliance for HAL 4.0 Rev 01.pdf | 200
16 | How are updates and revisions to technical documentation controlled? | document update and revision control procedure technical documentation version history and tracking standard operating procedure for documentation revisions | The control of updates and revisions to technical documentation for the Halcyon 4.0 radiotherapy system involves several processes and procedures as evidenced by different documents: 1. **Service Documentation Updates**: There are explicit instructions to review and update service documentation as needed. For example, the service instructions related to lubrication procedures and rail configurations are updated to ensure clarity ([Risk Management / Design FMEAs / D101531 - HAL CCH Rail Vendor Change – DFMEA Design Review.docx.pdf ¦ Page: 3]). 2. **Software Requirements Management**: The requirements management work instruction specifies the procedures for reviewing and updating requirements ([Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf ¦ Page: 28]). 3. **Configuration Management**: Change control after release is governed by processes in place, such as the DND01 Design and Development Procedure, which addresses the evaluation and implementation of upgrades. This control extends to third-party project components as mandated by Varian’s process for contractual change management ([Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf ¦ Pages: 61-62]). 4. **Approval and Traceability**: Before final approval, there are detailed checks and peer reviews to ensure the completeness and correctness of the updated documents. Traceability of the document updates and their approval is maintained, with the signed approval forms and documents filed in the Design History File ([Design Output Specification / Labeling / D65651 Halcyon-Ethos 4.0 IFU (P1058124-001-A) User Assistance Checklist Rev01.docx.pdf ¦ Pages: 4-5]). 5. **Formal Review and Validation**: There is a formal review process where the accuracy, clarity, and completeness of information are verified with reviewers, which helps in maintaining the integrity of technical documentation ([Design Output Specification / Labeling / D77711 Halcyon and Ethos Radiotherapy System 4.0 IFU (P1039628-003-C) Checklist Rev 01.docx.pdf ¦ Pages: 4-5]). 6. **Change Request and Maintenance Release**: Change requests and maintenance releases are systematically documented and reviewed, involving activities such as impact analysis and defining approval strategies ([Design Changes / Change Requests / D72997 HAL 4.0 MR1 Project Plan Change Request 1 Rev01.docx.pdf ¦ Pages: 12-13]). These documents collectively show a structured approach to managing updates and revisions to the technical documentation of the Halcyon 4.0 radiotherapy system. Documents cited: - Risk Management / Design FMEAs / D101531 - HAL CCH Rail Vendor Change – DFMEA Design Review.docx.pdf ¦ Page: 3 - Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf ¦ Page: 28, 61-62 - Design Output Specification / Labeling / D65651 Halcyon-Ethos 4.0 IFU (P1058124-001-A) User Assistance Checklist Rev01.docx.pdf ¦ Pages: 4-5 - Design Output Specification / Labeling / D77711 Halcyon and Ethos Radiotherapy System 4.0 IFU (P1039628-003-C) Checklist Rev 01.docx.pdf ¦ Pages: 4-5 - Design Changes / Change Requests / D72997 HAL 4.0 MR1 Project Plan Change Request 1 Rev01.docx.pdf ¦ Pages: 12-13 | Risk Management / Design FMEAs / D101531 - HAL CCH Rail Vendor Change – DFMEA Design Review.docx.pdf ¦ Page: 3 Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf ¦ Page: 28, 61-62 Design Output Specification / Labeling / D65651 Halcyon-Ethos 4.0 IFU (P1058124-001-A) User Assistance Checklist Rev01.docx.pdf ¦ Pages: 4-5 Design Output Specification / Labeling / D77711 Halcyon and Ethos Radiotherapy System 4.0 IFU (P1039628-003-C) Checklist Rev 01.docx.pdf ¦ Pages: 4-5 Design Changes / Change Requests / D72997 HAL 4.0 MR1 Project Plan Change Request 1 Rev01.docx.pdf ¦ Pages: 12-13 | 200
17 | Is traceability established from requirements to design to verification and validation? | document outlining traceability of requirements to design artifacts verification procedures linked to specific design requirements validation results showing requirements fulfillment and design integrity | The audit objective is to determine whether traceability is established from requirements to design to verification and validation for the Halcyon 4.0 radiotherapy system. 1. **Traceability from Requirements to Design to Verification:** - The "D68782 HAL 4.0 Requirement Tracing rev 01" document describes the upward tracing of requirements, indicating that requirements trace to risk control measures, standards, or user needs specific to Halcyon 4.0 [Design Output Specification / Design Trace Matrices / D68782 HAL 4.0 Requirement Tracing rev 01.pdf, Page 3]. - The "D68599 Hal 4.0 Design Verification Trace Matrix Rev 02" document confirms that there is a trace from requirements to test cases and that test results are stored within the document, providing traceability from requirements to verification activities [Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Page 3]. 2. **Verification and Validation Documentation:** - The IEC 62304+A1:2015 Software Lifecycle Assessment document states that traceability between software requirements and tests or other verification activities is recorded, showing that verification is linked to specified software requirements [Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, Page 47]. - Additional evidence of verification efforts, relevant procedures, and device records is present in several test reports and validation summaries which document test cases, pass/fail outcomes, and trace matrices relevant to the validation process [Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page 265-266]. 3. **Traceability to Validation:** - The "D68014 HAL 4.0 Validation Summary Report Rev 02" summarizes the results of the validation activities and refers to a validation trace matrix, indicating that tests conducted confirm user needs are met and are traced back to initial requirements [Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf, Page 3]. - The "D54550 HAL 4.0 Validation Plan Rev 05" outlines the validation activities, including a validation trace matrix which documents how validation efforts align with requirements and user needs [Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf, Page 18]. Overall, the documentation shows that traceability is indeed established from requirements to design, and from design to verification and validation activities for the Halcyon 4.0 system. Documents cited: - Design Output Specification / Design Trace Matrices / D68782 HAL 4.0 Requirement Tracing rev 01.pdf, Page 3 - Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Page 3 - Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, Page 47 - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page 265-266 - Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf, Page 3 - Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf, Page 18 | Design Output Specification / Design Trace Matrices / D68782 HAL 4.0 Requirement Tracing rev 01.pdf, Page 3 Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Page 3 Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, Page 47 Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page 265-266 Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf, Page 3 Design Validation / Verification and Validation Plans / D54550 HAL 4.0 Validation Plan Rev 05.docx.pdf, Page 18 | 200
18 | Does the technical documentation align with regulatory requirements of all applicable jurisdictions? | technical documentation compliance checklist for jurisdiction requirements comparison of device specification with jurisdictional regulatory standards review of documentation against specific regional regulatory codes and guidelines | The retrieved document excerpts provide a limited view of the technical documentation and its alignment with regulatory requirements for the Halcyon 4.0 radiotherapy system. However, there is some evidence indicating compliance with specific regulatory standards and checklist processes that take into account regulatory and standards content. 1. **Compliance with Regulatory Standards:** - The Halcyon 4.0 has been designed to meet international product safety standards. The specific standards mentioned include the EU Medical Device Regulation (MDR 2017/745), Health Canada's Medical Devices Regulations (CMDR SOR/98-282), and various FDA regulations in the United States such as 21 CFR parts 801, 803, 806, and 807 [Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, pages 30-31]. 2. **Checklists and Compliance Review:** - Several checklists (e.g., User Assistance Checklists) include sections that mention collaboration with content owners on required regulatory and standards content. These checklists have tasks to verify compliance with standards as part of their process [Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf, pages 1-6]. - Specific references indicate ensuring that documents meet regulatory requirements. For example, the application of Acrobat Actions for final distribution, verifying metadata per project name, product, and release version, and verifying technical compliance through DITA QA files [Design Output Specification / Labeling / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, pages 4-5]. 3. **Regulatory Plan for Specific Jurisdictions:** - The Halcyon 4.0 Regulatory Plan outlines the marketing clearance strategy for jurisdictions such as the EU, China, and others, indicating detailed processes for documentation requirements, product testing report submissions, and ensuring compliance with applicable standards [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 40-41]. 4. **Applicable Standards and Guidance Documents:** - The HAL 4.0 APS Record lists applicable standards and guidance documents, which support the compliance claims. It references various standards such as EN ISO 13485:2016, EN ISO 14971:2019, and others related to product compliance and quality assurance [Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, pages 7-19]. The current collected documents indicate that the Halcyon 4.0 technical documentation is structured to align with regulatory requirements, as evidenced through the application of checklists that incorporate regulatory content and plans that address jurisdiction-specific documentation needs. However, excerpts do not provide comprehensive coverage for all jurisdictions or all aspects of the regulatory requirements, and additional documentation might be required for a full audit. **Documents cited:** - Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, pages 30-31 - Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf, pages 1-6 - Design Output Specification / Labeling / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, pages 4-5 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 40-41 - Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, pages 7-19 | ** Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, pages 30-31 Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf, pages 1-6 Design Output Specification / Labeling / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, pages 4-5 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 40-41 Regulatory / Applicable Standards / D47457 HAL 4.0 APS Record Rev 04.docx.pdf, pages 7-19 | 200
19 | Are country-specific requirements addressed, such as FDA DHF, EU Technical File, or Health Canada STED? | FDA Design History File (DHF) compliance documentation EU Technical File requirements and contents Health Canada STED submissions and compliance requirements | Based on the retrieved document excerpts, the Halcyon 4.0 (HAL 4.0) radiotherapy system addresses multiple country-specific regulatory requirements: 1. **FDA Design History File (DHF):** - The FDA-related DHF documentation appears to be part of the overall design and development documentation, although specific details on FDA DHF requirements are not explicitly mentioned in the retrieved excerpts. 2. **EU Technical File:** - The HAL 4.0 complies with the EU Medical Device Regulation (MDR) 2017/745 and has specific documentation for EU requirements, including Medical Device Classification [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, pages: 53, 55]. - Specific documentation for EU marketing clearance strategy and ISD (International Submission Dossier) documentation requirements are outlined, indicating compliance with the EU's MDR [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages: 38, 42, 45]. 3. **Health Canada STED (Summary Technical Documentation for Demonstrating Conformity to the Essential Principles of Safety and Performance of Medical Devices):** - The HAL 4.0 complies with Health Canada Medical Device Regulations (CMDR SOR/98-282), and documentation is prepared for Health Canada Medical Device License Application [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, pages: 55; Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages: 46, 44-45]. - Specific reference to Canadian standards and documentation requirements indicate compliance with Health Canada's regulatory framework. While specific documentation regarding the precise contents of these submissions is mentioned, detailed contents or specific examples from the FDA DHF, EU Technical File, or Health Canada STED for the Halcyon 4.0 are not fully detailed in the provided excerpts. However, the existence of regulatory plans and certain documents indicates an attempt to meet these country-specific requirements. Documents cited: - Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, pages: 53, 55 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages: 38, 42, 44-46, 55 | Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, pages: 53, 55 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages: 38, 42, 44-46, 55 | 200
20 | How does the organization ensure updates to regulatory requirements are reflected in technical documentation? | technical documentation update process for regulatory compliance methods for tracking changes in regulatory requirements version control and history of technical documentation updates | The organization ensures updates to regulatory requirements are reflected in technical documentation through a structured approach as evident in the following excerpts: 1. **Compliance Updates**: The Halcyon 4.0 Regulatory Plan outlines several instances where compliance updates are required for various standards and documents. These requirements indicate that the organization tracks which documents need updates based on revised standards. For example, updates are required for compliance with standards like IEC 60601-1, IEC 60601-1-2, IEC 60601-2-68, and specific documentation such as the Summative Usability Test Report [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 40-41]. 2. **Document Update Responsibility**: The Regulatory Plan highlights responsibilities for updating specific documents, including the Statement of Indications for Use and Device Description to align with the latest regulatory requirements or templates [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, page 38]. This shows that specific roles within the organization are tasked with ensuring documents are updated as per current regulatory needs. 3. **Submission Documentation**: The plan specifies that all International Submission Dossier (ISD) documentation must be submitted electronically as they are approved and notes that additional documents may be required depending on evolving EU Regulatory requirements [01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, page 50]. 4. **Regular Monitoring and Adaptation**: The organization maintains a process for monitoring regulatory markets during projects and communicates changes to the team for adaptation of documentation, as shown in the Brazil Marketing Clearance Strategy [01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, page 63]. These processes collectively demonstrate that Varian Medical Systems uses its regulatory planning and management strategies to align technical documentation with updated regulatory requirements. However, the provided excerpts primarily detail the requirement and responsibility for updates rather than describing a complete process or a system in active place for accomplishing this routinely across all documentation. Documents cited: - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 38, 40-41 - 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, pages 50, 63 | Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, pages 38, 40-41 01. Design and Development Planning / 01. System / D30786 Halcyon 2.0 Regulatory Plan Rev 1.0.doc.pdf, pages 50, 63 | 200
21 | How is technical documentation controlled, including versioning, approvals, and access? | version control system for technical documentation approval workflow and authorization records for technical documentation access control measures for technical documentation | The control of technical documentation for the Halcyon 4.0 radiotherapy system includes aspects of versioning, approvals, and access as follows: 1. **Versioning and Documentation Control**: The Design History File Index for Halcyon 4.0 references all required project documentation with titles, identification numbers, and revision numbers. The documentation is stored electronically in the document control system "Nucleus," ensuring proper versioning and management [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14]. 2. **Approvals**: Approvals for user assistance and documentation involve electronic sign-offs, as evidenced by multiple documents that have been reviewed and electronically signed by designated approvers. For instance, document D68539, which deals with the build procedure, includes electronic signatures as a part of its approval process [Design Output Specification / Build Procedures / D68539 MPC 2.37 SBP Rev 01.pdf, Page 9]. Additionally, documents undergo a structured approval process where final distribution versions are proofread, approved, and signed off before being filed in the Design History File (DHF) [Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 3-4]. 3. **Access Control**: After approval, the finalized documents, such as user assistance plans, are uploaded to the GPL Customer Documentation system, where they are accessible to authorized personnel [Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 4-5]. Additionally, the document control system (Nucleus) ensures that only authenticated users have access to the critical documentation, further demonstrating control over who can access the information [Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14]. Documents cited: - Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14 - Design Output Specification / Build Procedures / D68539 MPC 2.37 SBP Rev 01.pdf, Page 9 - Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 3-4 - Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 4-5 | Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf, Page 13-14 Design Output Specification / Build Procedures / D68539 MPC 2.37 SBP Rev 01.pdf, Page 9 Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 3-4 Design Output Specification / Labeling / D96388 HAL 4.0 UA Legal Information (EU REP Change) UA Checklist Rev 01.docx.pdf, Page 4-5 | 200
22 | Who has responsibility for maintaining technical documentation? | roles and responsibilities for maintaining technical documentation department in charge of technical documentation updates staff assigned to oversee technical document management | The responsibility for maintaining technical documentation for the Halcyon 4.0 (HAL 4.0) radiotherapy system is distributed among several roles, with specific responsibilities detailed in various documents: 1. **User Assistance Developers (UADs)**: The User Assistance Plan (UAP) states that User Assistance Developers are part of the project team, attend project meetings, and report status regularly to the project team. They are responsible for ensuring the documentation is kept up-to-date throughout the project. The Project Roles, Team, and Responsibilities section further details that Technical Reviewers and Approvers are identified based on the project team defined in the DDP/SDP, and individuals are designated who are best suited to represent the required roles and responsibilities for any given publication [Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 3; Page: 7-8]. 2. **Product Marketing Manager**: Responsibilities of the Product Marketing Manager include providing product documentation for product submissions and registrations, as detailed in the responsibilities section of the regulatory plan [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 73]. 3. **Quality Assurance Oversight**: The Regulatory Plan document specifies that the Quality Assurance Lead provides oversight for maintaining quality within documentation related to registration and submissions [Design Output Specification / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, Page: 6]. 4. **Manufacturing Engineer**: For manufacturing documentation, the Manufacturing Engineer is tasked with performing the tasks described every time new versions of product software are released to manufacturing [Design Transfer / Manufacturing Documentation / D67017-Rev 02-HAL4.0 Golden Disk Build Instruction.pdf, Page: 2]. These excerpts indicate a decentralized approach where different roles within the organization are responsible for maintaining specific segments of the technical documentation, under a structured collaborative effort. However, it is not entirely clear who coordinates these responsibilities or how changes are managed across different segments of technical documentation in practice, indicating a potential gap in centralized oversight across all types of technical documentation for the HAL 4.0. Documents cited: - Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 3; Page: 7-8 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 73 - Design Output Specification / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, Page: 6 - Design Transfer / Manufacturing Documentation / D67017-Rev 02-HAL4.0 Golden Disk Build Instruction.pdf, Page: 2 | Design Output Specification / Labeling / D93257 Halcyon Evolve User Assistance Plan Rev02.docx.pdf, Page: 3; Page: 7-8 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, Page: 73 Design Output Specification / D84239 HAL 4.0 CRN UA Checklist Rev02.docx.pdf, Page: 6 Design Transfer / Manufacturing Documentation / D67017-Rev 02-HAL4.0 Golden Disk Build Instruction.pdf, Page: 2 | 200
23 | Is documentation readily retrievable for audit and regulatory submission? | process for retrieving audit documentation regulatory submission documentation access audit-ready document management system | Based on the provided excerpts, documentation for the Halcyon 4.0 radiotherapy system appears to be readily retrievable for audit and regulatory submission. Here is the supporting evidence: 1. **Design History File Documentation**: The Design History File (DHF) includes indexed documentation, such as verification test readiness reviews, design outputs, and risk management plans, showing organized record-keeping ([Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 30]). 2. **Retention and Accessibility**: The Varian System has the capability to store and manage audit logs with configurable retention policies, ensuring no entries are lost due to limited storage space and audit logs are protected from unauthorized access ([Design Transfer / Manufacturing Documentation / D81612-HAL 4U WS Automation Test Application Cybersecurity Summary Report- Rev01 - signed.pdf ¦ Page: 7-8]). 3. **Audit Capabilities**: The system supports logging various user interactions and security-related events, providing an audit trail for different user activities and changes within the system ([Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf ¦ Page: 52-53]). 4. **Approvals and Signatures**: Many documents include electronic signature records, ensuring traceability and approval by responsible personnel, which supports both legal and regulatory requirements ([Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf ¦ Page: 5-6]). 5. **Availability of Documentation for Submissions**: Regulatory plans like the HAL 4.0 Regulatory Plan mandate that all final documentation be electronically available before the submission deadlines, supporting readiness for regulatory submissions ([Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page: 30]). 6. **Electronic Documentation Archives**: Documents such as user assistance checklists and publication notices are archived and maintained in defined locations, indicating a structured process for document management ([Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf ¦ Page: 4-5]). 7. **Regulatory Compliance**: Documents necessary for regulatory compliance, like risk management reports and validation plans, are maintained in the DHF, demonstrating compliance with industry standards and requirements ([Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf ¦ Page: 4]) and ([Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf ¦ Page: 17-18]). These points collectively support the availability and organization of documentation necessary for audit and regulatory submission for the Halcyon 4.0 radiotherapy system. Documents cited: - Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 30 - Design Transfer / Manufacturing Documentation / D81612-HAL 4U WS Automation Test Application Cybersecurity Summary Report- Rev01 - signed.pdf ¦ Page: 7-8 - Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf ¦ Page: 52-53 - Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf ¦ Page: 5-6 - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page: 30 - Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf ¦ Page: 4-5 - Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf ¦ Page: 4 - Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf ¦ Page: 17-18 | Design Output Specification / D67731 HAL 4.0 Design History File Index Rev 05.docx.pdf ¦ Page: 30 Design Transfer / Manufacturing Documentation / D81612-HAL 4U WS Automation Test Application Cybersecurity Summary Report- Rev01 - signed.pdf ¦ Page: 7-8 Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf ¦ Page: 52-53 Design Output Specification / D84872 HAL 4.0 CRN UA Checklist.docx.pdf ¦ Page: 5-6 Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf ¦ Page: 30 Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf ¦ Page: 4-5 Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf ¦ Page: 4 Design Validation / Verification and Validation Plans / D70753 HAL 4.0 MR1 Validation Plan Rev 02.docx.pdf ¦ Page: 17-18 | 200
24 | Are document control processes followed? | document version control procedures document approval and distribution processes procedures for document review and update | The document control processes for Halcyon 4.0 are generally being followed, based on the evidence provided in the excerpts: 1. **Software Configuration Management and Change Control:** - The Halcyon 4.0 project follows a structured software configuration management process. Unique identification of configuration items and their versions is mandated by the DND19 Software Configuration Management Procedure, documented in the DND04T01 Design and Development Plan Rev 04 [Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 75]. - Changes to configuration items are controlled through an approved Change Request process mentioned in LL00000016 Change Request Rev 05, and implementation follows the established procedures in the Design and Development Procedure Rev 02 [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, page 299]. 2. **Approval Processes and Verification:** - Various documents related to Halcyon 4.0 have been reviewed and electronically signed by relevant participants, ensuring that document controls, such as versioning and approval, are in place. For instance, the "Process Control Plan for Packaging" document includes electronic signatures, indicating following of proper issuance and control processes [Design Transfer / Manufacturing Documentation / 70002131, Rev.0.1 - Process Control Plan for Packaging-HAL4.0 BASIC KV IMAGING KIT (Beijing) .pdf, page 4-6]. 3. **Archiving and Retention:** - The records retention policy is specified in DOC0203 Quality Records Retention Schedule Rev 07. This policy is referenced within the development procedure documents to ensure that documentation, testing, and other records are maintained as per regulatory requirements [Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 51]. 4. **Documentation of Changes and Results:** - Processes are established for reviewing and approving documents, which includes electronic signatures and maintaining this documentation in controlled versions, as exemplified by the risk management documentation [Risk Management / Risk Analyses / D60083 HAL 4.0 CTR Review Report rev01.pdf, page 10]. However, the provided excerpts do not include explicit evidence of a comprehensive audit process to confirm if all document control processes are followed across all areas, missing specific checks or audits of compliance beyond configuration and change management. Documents cited: - Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 75 - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, page 299 - Design Transfer / Manufacturing Documentation / 70002131, Rev.0.1 - Process Control Plan for Packaging-HAL4.0 BASIC KV IMAGING KIT (Beijing) .pdf, page 4-6 - Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 51 - Risk Management / Risk Analyses / D60083 HAL 4.0 CTR Review Report rev01.pdf, page 10 | Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 75 Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, page 299 Design Transfer / Manufacturing Documentation / 70002131, Rev.0.1 - Process Control Plan for Packaging-HAL4.0 BASIC KV IMAGING KIT (Beijing) .pdf, page 4-6 Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 51 Risk Management / Risk Analyses / D60083 HAL 4.0 CTR Review Report rev01.pdf, page 10 | 200
25 | Is each device uniquely identified and linked to its technical documentation? | unique identifier and serial number system for each device technical documentation linkage and reference guide device identification and documentation cross-reference matrix | The retrieved excerpts provide several pieces of evidence confirming that each device of the Halcyon 4.0 radiotherapy system is uniquely identified and linked to its technical documentation: 1. **Unique Device Identifier (UDI)**: - The Halcyon 4.0 is assigned a Basic UDI-DI for EU MDR compliance and specific Model UDIs. The UDI for Halcyon 4.0 is '4.000857235006075' [Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, Page: 4]. 2. **Identification through Serial Numbers**: - Device Serial Numbers are used and encoded as ASCII code representations for identification. An example provided is SN1004, encoded as '083078049048048052' [Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Pages: 293-294; Design Input Specification / System Requirement Specifications / D63662 HAL 4.0 System Requirements Specification rev 02.pdf, Page: 377]. - The labeling of devices includes essential information such as the serial number, which is present on the main device label along with compliance to ISO 15223-1:2016 standards [Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf, Page: 28]. 3. **Configuration and Version Identification**: - The Halcyon 4.0 system uses a configuration management process that establishes a scheme for unique identification of configuration items and their versions, ensuring traceability to technical documents [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages: 296-297; Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page: 284-285]. The excerpts demonstrate comprehensive documentation practices that uniquely identify each Halcyon 4.0 device through UDIs and serial numbers, linking these to technical specifications and compliance documents. Each device's UDI and serial number provide a direct link to its technical documentation, ensuring traceability throughout its documentation lifecycle. Documents cited: - Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, Page: 4. - Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Pages: 293-294, 722. - Design Input Specification / System Requirement Specifications / D63662 HAL 4.0 System Requirements Specification rev 02.pdf, Page: 377. - Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf, Page: 28. - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages: 296-297. - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page: 284-285. | Design Output Specification / Design Descriptions / D58308 HAL 4.0 Device Description Rev 1.0.docx.pdf, Page: 4. Design Output Specification / Design Trace Matrices / D68599 Hal 4.0 Design Verification Trace Matrix Rev 02.pdf, Pages: 293-294, 722. Design Input Specification / System Requirement Specifications / D63662 HAL 4.0 System Requirements Specification rev 02.pdf, Page: 377. Regulatory / Essential Requirements Checklist / D70060 HAL 4.0-GSPR Checklist.pdf, Page: 28. Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, Pages: 296-297. Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, Page: 284-285. | 200
26 | How are product variants and configurations managed within technical documentation? | product variant management procedures in technical documentation configuration management protocols for technical documents documentation guidelines for handling product variants and configurations | Based on the excerpts provided, product variants and configurations for the HAL 4.0 device are primarily managed through software configuration management and regulatory documentation. Here's how they are managed: 1. **Configuration Management:** The DND19 Software Configuration Management Procedure is a critical document that outlines the management of configuration items and their versions for Halcyon 4.0. It addresses the identification, documentation, change control, and overall management of configuration items, which are essential for maintaining different product configurations and variants [Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 18-19, 51; Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf, page 231-232, 286-287]. 2. **Design Development Plan (DDP):** The DND04T01 Design and Development Plan (DDP) includes details on configuration management and references procedures for managing software configurations. This ensures that different variations or configurations of the HAL 4.0 system are controlled and documented appropriately [Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf, page 76-77]. 3. **Change and Version Control:** Changes to configuration items are controlled through approved change requests as documented in the DND04F01 Requirements or Project Plan Change Request. This aspect helps accommodate product variants by managing changes systematically and maintaining version control [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf, page 240-242]. 4. **Project Documentation:** Regulatory documentation such as the HAL 4.0 Regulatory Plan Rev 5.0 includes ISD documentation requirements, which are essential for maintaining records of product variants and configurations for regulatory submissions across different jurisdictions. The documentation updates and manages the inclusion of different product variants and any related technical changes [Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf, page 37-44]. 5. **Software and Documentation Updates:** The software versioning and documentation (such as the Instructions for Use, Quick Reference Guides) are updated and managed to reflect different configurations and variants of the Halcyon system, ensuring that documentation is accurate and complete for the specific variant or configuration [Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf, page 2-3]. These processes, embedded within the documented procedures and plans, ensure that product variants and configurations of the HAL 4.0 system are managed methodically in the technical documentation. Documents cited: - Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf - Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf - Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf - Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf | Regulatory / Applicable Standards / D68985 HAL 4.0 IEC 62304+A1 2015 Software Lifecycle Assessment Rev 01.docx.pdf Regulatory / Test Reports - Regulatory Compliance / D101277 106485676MIN-001 IEC60601-1 Ed.3.2 HAL4.0_30_APR_2026 Rev 1.pdf Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf Regulatory / Regulatory Plans / D54978 HAL 4.0 Regulatory Plan Rev 5.0.pdf Design Output Specification / Labeling / D63983 Halcyon 4.0 UA Checklist Rev 01.pdf | 200
27 | Is traceability maintained across product families and versions? | product family traceability documentation version control and traceability records traceability matrix for product versions and families | The evidence suggests that traceability is maintained across product families and versions for the Halcyon 4.0 radiotherapy system. The traceability is documented in several excerpts, as follows: 1. **Traceability Plan**: The document [Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf] outlines the traceability plan for the HAL4.0 product platform, which includes the Halcyon 4.0 and Ethos Radiotherapy System 4.0 [source, page 4]. This document specifies that all medical device components must adhere to traceability requirements, with any future changes necessitating an update to the plan [source, page 4]. Additionally, traceability is evaluated and approved by various departments, ensuring comprehensive oversight [source, pages 5-6]. 2. **Verification and Validation Summary**: Various validation summary reports detail how traceability is recorded, including mapping requirements to validation test outcomes. For example, the [Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf] advises referring to the Validation Trace Matrix for validation test results and associated user needs and risk control measures [source, page 11]. 3. **Design and Development Documentation**: The system's design and development documentation, including pervasive references to a "Design Trace Matrix," ensures that all requirements are traceable through the document matrix, offering a structured path from requirements, to implementation, to testing [Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf ¦ page 293]. 4. **Risk Management Documentation**: The HAL 4.0 Risk Control Measures Trace Matrix [source: Risk Management / Risk Management Reports / D68779 HAL 4.0 Risk Control Measures Trace Matrix rev01.pdf] further supports the traceability framework by cross-referencing verification measures and risk control measures [source, many pages]. 5. **Cybersecurity Traceability Analysis**: The [Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf] also contributes to the traceability effort by mapping cybersecurity requirements and confirming their addressal through risk management processes [source, page 8]. In each of these documents, there is a structured approach to maintaining traceability across different components and iterations of the HAL4.0 system, corroborating that traceability is maintained for the product line. These approaches include detailed process documentation, oversight from various departments, and consistent updates when changes occur, ensuring a robust traceability mechanism across product families and versions. Documents cited: - Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf - Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf - Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf - Risk Management / Risk Management Reports / D68779 HAL 4.0 Risk Control Measures Trace Matrix rev01.pdf - Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf | Design Transfer / Manufacturing Documentation / D58685-Rev04-HAL4.0 Traceability Plan.pdf Design Validation / Verification and Validation Summary Reports / D68014 HAL 4.0 Validation Summary Report Rev 02.docx.pdf Regulatory / Test Reports - Regulatory Compliance / 104993953MIN-001 - HAL4.0 - IEC 60601-1.pdf Risk Management / Risk Management Reports / D68779 HAL 4.0 Risk Control Measures Trace Matrix rev01.pdf Risk Management / Risk Management Reports / D68683 HAL 4.0 Cybersecurity Traceability Analysis Rev 01.pdf | 200